## Slide 1: ICRP Publication 121
Radiological Protection in Paediatric Diagnostic and Interventional Radiology
Authors on behalf of ICRP
P-L. Khong, H. Ringertz, V. Donoghue, D. Frush, M. Rehani, K. Applegate, R. Sanchez

## Slide 2: Increased risk in Infants and Children
Diagnostic radiological examinations carry higher risk per unit of radiation dose, on average, for the development of cancer in infants and children compared to adults.
The higher risk is explained by the longer life expectancy in children for any harmful effects of radiation to manifest and the fact that developing organs and tissues are more sensitive to the effects of radiation.
Moreover, the average risk is higher in infants and young children compared with older children
ICRP 121, Introduction, Pg 15

## Slide 3: General Aspects of Radiological Protection
Justification
  the use of the radiological examination in question will do more good than harm to the patient
  the specific radiological examination when required for a specific disease and age group has a specified objective and this will usually improve the diagnosis or treatment or will provide necessary information about the exposed individuals
  the examination is required for that individual patient.
Optimization
  ALARA
  Radiological equipment for children
  Adequacy of equipment and technical parameters
  Diagnostic reference level (DRL)
Quality criteria implementation and audit
ICRP Publication 103

## Slide 4: Justification
Imaging procedure is reliable, i.e. reproducible, and of sufficient sensitivity, specificity, accuracy
Radiologist able to make accurate interpretation
Performed by qualified radiographer/technologist with appropriate monitoring for quality and safety by medical physicists
Accountability by single person, usually radiologist
Imaging techniques that do not employ the use of ionizing radiation should always be considered as an possible alternative.

## Slide 5: Some examples of radiographic examinations not routinely justified
  Skull radiograph in an infant or child with epilepsy
  Skull radiograph in an infant or child with headaches
  Sinus radiograph in an infant or child under 6 years suspected of having sinusitis
  Cervical spine radiograph in an infant or child with torticollis without trauma
  Radiographs of the opposite side for comparison in limb injury
  Scaphoid radiographs in children under 6 years
  Nasal bone radiographs in children under 3 years
ICRP 121, Pg 22

## Slide 6: Optimisation: ALARA
Optimisation of radiological protection in such a way that the required image is obtained with least radiation dose and net benefit is maximised
ALARA (as low as reasonably achievable) principle should be adhered to for every examination

## Slide 7: Radiological Equipment and adequacy of radiological equipment and technical parameters
At procurement, installation and commissioning
Radiographic equipment should have broadest range of settings to optimise the dose to the size of the child
Modifications of parameters may be necessary both at installation and later in the use of the equipment
ICRP 121, Pg 24

## Slide 8
Special consideration to availability of dose reduction methods e.g.
  Virtual collimation
  Low attenuation table tops
  Removable grids
  Pulsed fluoroscopy and last image hold
  Spectral filters and adaptive technologies to minimise blooming
  Adding a copper filter in addition to the aluminium filtration should be considered
Conventional Radiography and Fluoroscopy: Equipment

## Slide 9: Diagnostic reference levels (DRLs)
Assist in optimisation process
If value is exceeded regularly, practice should be investigated.
The upper DRL is often taken as the third quartile value, i.e. the value below which the measurements for three quarters of the institutions lie.
Measurements taken in any institution should lie below the upper DRL, and if above, it should be possible to reduce exposures below the DRL without loss of clinical information.
ICRP Publication 103

## Slide 10
Examples of Diagnostic Reference Levels for standard five-year-old patients, expressed in entrance surface dose per image for single views. | 
Radiograph | 5-year-old patients Entrance surface dose Per single view (mGy)*
Chest Posterior Anterior (PA) | 0.1
Chest Anterior Posterior (AP for non co-operative patients) | 0.1
Chest Lateral (Lat) | 0.2
Skull Posterior Anterior/Anterior Posterior (PA/AP) | 1.5
Skull Lateral (Lat) | 1.0
Pelvis Anterior Posterior (AP) | 0.9
Abdomen (AP/PA with vertical/horizontal beam) | 1.0
*Upper DRL expressed as entrance surface dose to the patient. The entrance surface dose for standard-sized patients is the absorbed dose in air (mGy) at the point of intersection of the beam axis with the surface of a paediatric patient, backscatter radiation included. | 
European Commission 1996

## Slide 11: Quality criteria implementation and audit
Need for follow-up and regular audits after implementation of quality criteria
Audits of referral criteria, image quality, and imaging technique in paediatric radiology practices have found that better results are obtained for paediatric specialist centers compared with non-specialist centers. Thus, sharing of good practice by paediatric specialist centers is important for improving practices and patient outcomes.
ICRP 121, Pg 27

## Slide 12: Conventional Radiography and Fluoroscopy
Proper patient positioning and applying immobilization which is often required in infants and young children
Appropriate field size and correct X-ray beam limitation
Use collimation to expose only the area intended for examination
ICRP 121, Pg 29

## Slide 13: Protective Shielding
Breasts, gonads, and/or thyroid < 5cm to the primary beam should be protected whenever possible without impairing the necessary diagnostic information
Male gonads
    Not to be included within the primary radiation field
Female gonads
    Gonad protection may not be possible; e.g. for the indications of trauma, incontinence, abdominal pain, misplaced shielding may mask important pathology.
Lens
    SXR: PA rather than AP
Breasts
    CXR and Spine X-Ray: PA rather than AP

## Slide 14: Radiographic exposure conditions
Knowledge and correct use of appropriate radiographic factors e.g. focal spot size, filtration, focus to image plane distance, tube current-exposure time product
Exposure times
  Short exposure times required due to motion
Higher speed classes of screen film system
  Peripheral skeleton; speed class of 200-400
  Others; speed classes of 400-800
ICRP 121, Pg 31

## Slide 15: Radiographic exposure conditions
Additional filtration
Anti-scatter grid
  In neonates, infants and younger children, anti-scatter grid often not necessary because of relatively low scatter radiation
  Removable grids for fluoroscopy
Automatic exposure control (AEC)
  Due to the wide range in size of paediatric patients, there is need to optimise AEC devices for handling patients accordingly
Exposure time
  Exposure times should be short because children do not generally cooperate and are difficult to restrain, so equipment should work and provide reproducibility in the shortest time range

## Slide 16: Fluoroscopy
Grid-controlled pulsed fluoroscopy systems
Keep fluoroscopy table as far from X-ray source as possible and image intensifier as close to patient as possible
Minimize scatter radiation with hanging lead drapes
Use pulsed fluoroscopy for monitoring procedures (3-8 pulses/sec adequate)
Image intensifier positioned over area of interest before fluoroscopy is commenced
Use virtual collimation to help positioning
Apply tight collimation
ICRP 121, Pg 36

## Slide 17: Fluoroscopy
Angle beam away from radiosensitive areas
Limit use of magnification
Use still images from last-image hold or archive digital images to review findings
Consider alarm bells for fluoroscopy beyond a certain time
Record radiation dose (air kerma-area-product)
AAPM, 1998

## Slide 18: Interventional Radiology
Increasing complexity and length of procedures, thus overall radiation dose to the patient may be greater.
Increased dose to patients due to large size of image intensifier compared to patient and increased need to use magnification (must use collimation).
ICRP 121, Pg 39

## Slide 19: Interventional Radiology
A second, specific level of training in radiation protection is desirable, and in some countries mandatory.
Major paediatric interventional procedures, particularly in small infants, should be performed by experienced paediatric interventional operators.
Consider ultrasound guidance whenever possible
ICRP Publication 85

## Slide 20: Reducing unnecessary dose to staff
Paediatric procedures may take a longer time
Frequently necessary to come close to or enter the beam due to small size of child
Angulation of beam off the hands and apply strict collimation
Use protective devices (however, caution with gloves as it can reduce dexterity)
Power injector instead of hand injecting contrast material
ICRP 121, Pg 39

## Slide 21: Image acquisition during DSA
Each run should be necessary
Use fewest number of frames/sec
Images obtained using lower magnification (magnify using post-processing instead)
Tight collimation to include only area of interest
Use last image hold, image capture and archiving of runs
ICRP 121, Pg 41

## Slide 22: CT: Justification
Imaging techniques that do not employ the use of ionizing radiation should always be considered as an possible alternative.
Especially in children with chronic illness who require repeated imaging evaluation.
Ultrasound should be first line imaging modality for the abdomen e.g. acute appendicitis, intussusception
MRI is often best for evaluation of detailed information of soft tissues, nervous system or bone marrow
ICRP 121, Pg 46

## Slide 23: CT: Justification
Chronic illness and malignancy:
  Follow up CT scans should not be done not too early
  Alternative modalities should be considered for follow up
  Consider reducing CT scan volume and adjusting parameters
Multi-phase CT scans:
  Justification for every additional phase needed
ICRP 121, Pg 46

## Slide 24: Optimisation: CT Equipment
Special consideration should be given to dose reduction measures when purchasing new equipment for paediatric use
Medical physicist advice for procurement, commissioning, quality control tests etc.
ICRP 121, Pg 47

## Slide 25: CT software and hardware for dose reduction
Tube current modulation and automatic exposure control (AEC): ‘one-size does not fit all’
Organ-based dose modulation
Auto kV technology
X-ray detectors of improved efficiency
Iterative reconstruction
Methods to block unnecessary radiation from ‘helical overdosing’; dynamic or adaptive collimation
Filters for control of irradiation beam e.g. bow-tie filters
ICRP 121, Pg 48

## Slide 26: Optimisation: adjustment of scan parameters
Image quality: noise vs contrast
Adjustment of scan parameters (mAs, kVp, pitch) to weight or age
Use of automatic exposure control (AEC) techniques
  Noisier images, if sufficient for radiological diagnosis, should be accepted

## Slide 27: Image quality and Study quality
Depends on structure and region examined
  More noise acceptable in skeleton or lung compared to brain and abdominal examinations
  Tube current-time product (mAs) can be significantly reduced in children .e.g. 50mAs for CT chest
  80kVp suggested for infants under 5kg
Depends on clinical indication for the study
  More noise maybe tolerated for follow-up study

## Slide 28: Study Quality and Display
Motion artefact
    Selective use of sedation
    Adjustment of scan time and pitch, and dual-source CT
Meticulous use of IV contrast
Post-processing: planar reformation and 3D reconstruction
Display of images
  Optimise monitors
  Ambient environment

## Slide 29
Country-wide doses for paediatric CT head, chest and abdomen/pelvis
CT head |  |  |  |  |  | 
 | b1 (or 0-1) yrs |  | 5 (or 2-5) yrs |  | 10 (or 6-10) yrs | 
 | aCTDIvol 16 | DLP 16 | CTDIvol 16 | DLP 16 | CTDIvol 16 | DLP 16
UK 2005 | 35/30 | 270 | 50/45 | 470 | 65/50 | 620
Germany 2008 | 33 | 390 | 40 | 520 | 50 | 710
cSwitzerland 2008 | 20 | 270 | 30 | 420 | 40 | 560
France 2009 | 30 | 420 | 40 | 600 | 50 | 900
Greece 2009 | - | - | - | 650 | - | 975
Belgium 2010 | 35 | 280 | 43 | 473 | 49 | 637
CT chest |  |  |  |  |  | 
 | 1 (or 0-1) yrs |  | 5 (or 2-5) yrs |  | 10 (or 6-10) yrs | 
 | aCTDIvol 32 (16) | DLP 32 (16) | CTDIvol 32 (16) | DLP 32 (16) | CTDIvol 32 (16) | DLP 32 (16)
UK 2005 | 6 (12) | 100 (200) | 6.5 (13) | 115 (230) | 10 (20) | 185 (370)
Germany 2008 | 1.7 (3.5) | 28 (55) | 2.7 (5.5) | 55 (110) | 4.3 (8.5) | 105 (210)
cSwitzerland 2008 | 2.5 (5) | 55 (110) | 4 (8) | 100 (200) | 5 (10) | 110 (220)
France 2009 | 3 (6) | 30 (60) | 3.5 (7) | 63 (126) | 5.5 (11) | 137 (274)
Greece 2009 | - | - | - | 168 (336) | - | 289 (578)
Belgium 2010 | 4.2 (8.4) | 38 (76) | 4.7 (9.3) | 55.5 (111) | 4.5 (9) | 72 (144)
dUSA 2008 | 4.3 (8.5) | - | 4.8 (9.5) | - | 5.5 (11) | -
CT abdomen/pelvis |  |  |  |  |  | 
 | 1 (or 0-1) yrs |  | 5 (or 2-5) yrs |  | 10 (or 6-10) yrs | 
 | aCTDIvol 32 (16) | DLP 32 (16) | CTDIvol 32 (16) | DLP 32 (16) | CTDIvol 32 (16) | DLP 32 (16)
UK 2005 | - | - | - | - | - | -
Germany 2008 | 2.5 (5) | 70 (145) | 4 (8) | 125 (255) | 6.5 (13) | 240 (475)
cSwitzerland 2008 | 3.5 (7) | 65 (130) | 4.5 (9) | 150 (300) | 6.5 (13) | 190 (380)
France 2009 | 4 (8) | 80 (160) | 4.5 (9) | 121 (242) | 7 (14) | 245 (490)
Greece 2009 | - | - | - | 420 (840) | - | 560(1120)
Belgium 2010 | 3.9 (7.8) | 50.2 (101) | 5.5 (11) | 104.5 (209) | 4.8 (9.5) | 119 (238)
dUSA 2008 | 4.3 (8.5) | - | 5.0 (10) | - | 5.5 (11) | -
ICRP 121, Pg 44

## Slide 30: References
aFor head CT, CTDI and DLP values refer to 16-cm phantom. For chest and abdomen/pelvis CT, values refer to the 32-cm phantom, followed by the corresponding 16-cm phantom value in parentheses. Data have been adapted from the original publications, expressed according to the 16-cm phantom [1, 3, 5], the 32-cm phantom [4] or both [2].
bProposed DRLs expressed for children ages 1, 5 and 10 years [1, 4, 5] or using age ranges [2, 3]. Most pediatric DRL surveys do not include a specific 15-year-old category, although some include an 11- to 15-year-old group [2, 3]; the adult DRL in that country, or a value intermediate between adult and 10-year-old DRL may be considered appropriate for teenagers.
cSwitzerland subsequently adopted the values from the larger German [2] study.
dValues calculated according to recommendations of the Alliance for Radiation Safety in Pediatric Imaging, based on the future French DRL values for adult abdominal CT recommended by IRSN in 2008.
References:
Shrimpton, P.C., Hillier, M.S., et al., 2005. Doses from computed tomography (CT) examinations in the UK – 2003 review (NRPB-67). Available via www.hpa.org.uk/radiation.publication/index.htm. Accessed 10 Jan 2011.
Galanski, M., Nagal, H.D., Stamm, G. 2007. Paediatric CT exposure practice in the federal republic of Germany. Results of a nationwide survey in 2005/6. Medizinishe Hochschule Hannover. Available via www.mh-hannover.de/fileadmin/kliniken/diagnostische_radiologie/download/Report_German_Paed-CT Survey_2--5_06. pdf. Accessed 9 Jan 2011.
Verdun, F.R., Gutierrez, D., Vader, J.P., 2008. CT radiation dose in children: a survey to establish age-based diagnostic reference levels in Switzerland. Eur Radiol 18, 1980-1986.
Brisse, H.J., Aubert, B., 2009. CT exposure from pediatric MDCT: results from the 2007-2008 SFIPP/ISRN survey. J Radiol 90, 207-215.
Yakoumakis, E., Karlatira, M., Gialousis, G., et al., 2009. Effective dose variation in pediatric computed tomography : dose reference levels in Greece. Health Phys 97, 595-603.
Buls, N., Bosmans, H., Mommaert, C., et al., 2010. CT paediatric doses in Belgium: a multi-center study: results of a dosimetry audit 2007-2009. Available via http://www.fanc.fgov.be/CWS/GED/pop_View.aspx?LG=1&ID=2449

## Slide 31: Strategies to reduce CT radiation dose in Children
Rigorous justification of examinations
Prepare the patient
Acceptance of images with greater noise if diagnostic information can be obtained
Optimization of scan protocols
Limit scan coverage
Reduction of repeated scanning of identical area
ICRP 121, Pg 51

## Slide 32: Protective Shielding
Controversial
Protective shielding
  Breast bismuth protection
  Eye lens
  Thyroid
  Gonads
Bismuth protection should only be placed after scout view (or AEC pre-scanning)
ICRP 121, Pg 50

## Slide 33: Protective Shielding
Shielding may increase radiation dose or affect image quality if not placed optimally
AEC and dose-modulation systems increasingly available, and beam shielding interferes with these systems
Use of proper field size limitation and appropriate parameters
ICRP 121, Pg 50

## Slide 34: Protective Shielding
Studies on paediatric patients and phantoms have shown dose reduction of up to >50%
Suggested techniques to reduce streak artifacts:
  Shield needs to be appropriately placed with enough distance, > 2cm offset
  Smoothly placed over the surface
Protocols should be tested specifically for the scanner

## Slide 35: Summary and Recommendations
Justification of every examination involving ionising radiation, followed by optimisation of radiological protection is important in every patient, and especially in paediatric patients in view of the higher risk of adverse effects per unit of radiation dose compared to adults.
According to the justification principle, if a diagnostic imaging examination is indicated and justified, this implies that the risk to the patient of not doing the examination is greater than the risk of potential radiation induced harm to the patient.
Imaging techniques that do not employ the use of ionising radiation should always be considered as a possible alternative.
Optimisation of radiological protection involves optimised functioning of radiological equipment and quality control, ensuring radiological equipment and technical parameters are adequately tailored for paediatric patients and the implementation of diagnostic reference levels (DRL) to assist in the optimisation process.
Quality criteria implementation and regular audits should be instituted as part of the radiological protection culture in the institution.

## Slide 36
Attention should be paid to good radiographic technique including positioning and immobilisation of paediatric patients, field size and protective shielding, and radiographic exposure parameters should be specially tailored for patient size and age.
As most imaging equipment and vendor specified protocols are often structured for adults, modifications of equipment and exposure parameters may be necessary for paediatric use. Advice of medical physicists should be sought if possible to assist with installation, setting imaging protocols and optimisation.
Interventional procedures should be performed by experienced paediatric interventional staff due to the potential for high patient radiation dose exposure, and additional training in radiological protection to protect both patient and staff is recommended.
For CT, dose reduction should be optimised by adjustment of scan parameters (mA, kVp and pitch) according to patient weight or age, and weight-adapted CT protocols have been suggested and published. For the purpose of minimising radiation exposure, noisier images, if sufficient for radiological diagnosis, should be accepted. Optimised study quality also depends on region scanned and study indication. Other dose reduction strategies include restricting multiphase examination protocols, avoiding overlapping of scan regions, and scanning only the area in question. Furthermore, study quality may be improved by image post-processing to facilitate radiological diagnoses and interpretation.